# How does AI-driven parametric personalization using anthropometric data improve the functional fit and user comfort of 3D-printed upper limb prosthetics compared to traditional sizing methods?

AI-driven parametric personalization using anthropometric data enables better prosthesis-anatomy matching, resulting in improved functional fit with surface deviations under 1 mm and enhanced user comfort compared to conventional sizing methods.

## Abstract

AI-driven  parametric  design  using  anthropometric  data  enables  a  closer  match  between  prosthesis  geometry  and  individual  anatomy.  Studies  employing  high-fidelity  3D  scans  or  accessible  photo‐based  measurements  report  that  customized  sockets  and  components  yield  fit  improvements  such  as  surface  deviations  under  1 mm  and  increased  design  adaptability.  In  one  evaluation,  a  personalized  myoelectric  prosthesis  attained  a  mean  functional  performance  score  of  7.4  out  of  10  (BAM-ULA),  while  another  noted  that  QUEST  2.0  feedback  from  a  single  user  favored  a  3D‐printed  design  over  conventional  devices.

Several  reports  indicate  that  these  digital  workflows  reduce  fabrication  time  and  cost,  thereby  enhancing  user  comfort  and  reducing  muscle  fatigue  compared  to  traditionally  sized  devices.  A  subset  of  studies  applied  machine  learning  to  optimize  socket  structure  or  to  classify  EMG  signals,  suggesting  an  emerging  role  for  AI  in  refining  design  parameters.  Overall,  the  evidence—though  largely  derived  from  small,  exploratory  evaluations—supports  that  AI-driven  personalization  can  improve  functional  fit  and  user  comfort  relative  to  conventional  sizing  methods.

## Paper search

We performed a semantic search using the query "How does AI-driven parametric personalization using anthropometric data improve the functional fit and user comfort of 3D-printed upper limb prosthetics compared to traditional sizing methods?" across over 138 million academic papers from the Elicit search engine, which includes all of Semantic Scholar and OpenAlex.

We retrieved the 494 papers most relevant to the query.

59 sources were uploaded to Elicit for screening.

## Screening

We screened in sources that met these criteria:

- Prevalence and incidence of amputation: Does the paper present data or statistics about limb amputations (global, European, or national)?

- **Etiological factors of amputation **: Does it discuss congenital, traumatic, vascular, or infectious causes of amputation?

- Physical and emotional impacts of amputation: Does it analyze functional, psychological, or social consequences of amputation?

- Rehabilitation and clinical follow-up: Does it describe rehabilitation methods or clinical support for amputees?

- Function and biomechanics of the human hand: Does it address hand anatomy, motion, or functional complexity?

- Historical evolution of prosthetics: Does it present the historical or technological evolution of prosthetic devices?

- Classification of prosthetic devices: Does it classify or describe different prosthetic types (mechanical, myoelectric, passive, hybrid)?

- Challenges in using traditional prostheses: Does it discuss comfort, aesthetics, cost, or reasons for prosthesis abandonment?

- Industrial design and health design: Does it explore the role of industrial design in healthcare or assistive product development?

- User-centered or ergonomic approach: Does it mention usability, ergonomics, or participatory design with users?

- Digital fabrication and 3D printing: Does it involve additive manufacturing, digital modeling, or 3D printing for prosthetics?

- Parametric or AI-based modeling: Does it use parametric design, algorithms, or artificial intelligence for product generation?

- Anthropometric data and personalization: Does it use anthropometric data or focus on custom fit and morphology adaptation?

- Personalization and customization of prostheses: Does it address individualized design, comfort, or user-specific adaptation?

- Accessibility and equity of access: Does it discuss affordability, accessibility, or inequality in access to assistive devices?

- Open-source or collaborative approach: Does it promote open-source sharing, digital collaboration, or distributed fabrication?

- Ethical and social impact of assistive technology: Does it reflect on ethics, inclusion, or social sustainability of prosthetic design?

- Experimental validation or usability testing: Does the study include prototype testing or user evaluation?

- Publication year ≥ 2017: Was the study published in or after 2017?

- Peer-reviewed publication: Was it published in a peer-reviewed journal or academic conference?

We considered all screening questions together and made a holistic judgement about whether to screen in each paper.

## Data extraction

We asked a large language model to extract each data column below from each paper. We gave the model the extraction instructions shown below for each column.

- Research Objective:

- Extract the main research goal or purpose of the paper, including:

- The core problem the study addresses

- The specific aim or hypothesis

- Whether the study focuses on prosthetic design, usability, digital fabrication, or AI-driven systems

- Technology / Approach Used:

- Extract details about the technology or methodological approach described, including:

- Type of technology (3D printing, CAD, AI, parametric modeling, sensors, etc.)

- Application context (prosthetic design, health, rehabilitation, assistive devices)

- Level of technological maturity (concept, prototype, validated product)

- Personalization Methods:

- Extract details about the AI-driven or parametric personalization approach including:

- Specific AI/machine learning methods used (algorithms, models)

- Types of anthropometric data collected (measurements, scanning methods, body parameters)

- How anthropometric data was processed or integrated into design

- Parametric modeling or customization frameworks

- Data sources used for personalization (3D scans, manual measurements, databases)

- Any automation or optimization algorithms employed

- Comparison / Baseline Methods:

- Describe the traditional or baseline methods used for comparison, including:

- Standard sizing, manual fitting, or conventional prosthetic design workflows

- How body measurements were collected and applied

- Reference systems or manufacturing processes used as control

- if no direct comparison was made, note the absence or type of baseline used

- Materials and Fabrication:

- Extract structured details about materials and fabrication methods mentioned in the study, including:

- Material name (e.g., PLA, ABS, resin, metal alloys, silicone)

- Fabrication method (e.g., 3D printing, FDM, SLA, injection molding, CNC)

- Printer or equipment used (if specified)

- Mechanical or surface properties (strength, flexibility, finish)

- Customization details (e.g., modularity, multi-material design)

- Post-processing or finishing techniques

- Usability / Ergonomic Evaluation:

- Extract information about usability testing and ergonomic assessment, including:

- Type of evaluation (user testing, comfort analysis, feedback sessions)

- Number or type of participants involved

- Metrics used (usability scales, satisfaction, precision, comfort)

- Reported qualitative or quantitative results

- Accessibility and Cost:

- Extract any mention of accessibility, affordability, or open-source availability, including:

- Estimated cost or price reduction compared to conventional devices

- Mention of open-source sharing or digital distribution

- Barriers to access or strategies to improve affordability

- Ethical and Social Impact:

- Extract references to ethical, social, or inclusion aspects, such as:

- Discussion of equity, accessibility, or social responsibility

- Mentions of sustainability or user empowerment

- Broader impact on healthcare access or inclusivity

- Limitations and Challenges:

- List the main limitations or challenges identified by the authors, including:

- Technical constraints (hardware, accuracy, materials)

- Data or model limitations

- User adoption or accessibility issues

- Suggestions for improvement

- Future Work / Research Opportunities:

- Extract any mention of proposed future directions, improvements, or open questions, including:

- Plans for validation or user testing

- Integration with new technologies (AI, robotics, sensors)

- Expansion of datasets or design frameworks

## Report

Due to the limitations of the AI model, we are only able to process 40 sources while writing a report. This report was written using the 40 sources that had the highest screening scores out of the 188 sources that we screened in and extracted data from.

# Results

## Characteristics of Included Studies

The  40  included  records  comprise  a  mix  of  primary  research  studies,  design  and  development  reports,  and  literature  reviews  focused  on  applying  digital  fabrication  and  personalization  to  upper  limb  prosthetics.  A  significant  limitation  of  this  review  is  that  a  majority  of  the  included  records  were  available  as  abstracts  only,  which  prevented  a  full  assessment  of  their  methodology  and  findings.  Furthermore,  several  records  appear  to  be  duplicates  based  on  identical  or  near-identical  abstract  content  (e.g.,  P1/P38/P39;  P3/P10),  which  may  artificially  inflate  the  volume  of  available  evidence.

The  primary  studies  are  predominantly  preliminary  or  exploratory,  often  featuring  small  sample  sizes  (frequently  n=1  to  n=10)  or  case-study  designs.  The  main  focus  across  most  studies  is  on  the  design,  development,  and  fabrication  of  customized  prostheses,  with  a  strong  emphasis  on  leveraging  3D  printing  and  parametric  design  to  enhance  affordability  and  accessibility.  While  many  studies  report  on  usability  or  functional  outcomes,  these  are  typically  secondary  to  the  description  of  the  technology  itself  and  often  lack  rigorous,  comparative  evaluation  against  traditional  methods.

| Study ID | Study Focus | Methodology Type | Prosthetic Component | Primary Outcomes | Full text retrieved |
| --- | --- | --- | --- | --- | --- |
| P1 | Parametric design for customized transradial sockets | Design & Development Report | Socket | Design time optimization, preliminary user fit | Yes |
| P2 | Data-driven framework for optimizing socket adaptability | Primary Study (n=5) | Socket | Improved adaptability, reduced muscle fatigue | Yes |
| P3 | AI-driven personalized myoelectric prosthesis | Primary Study (n=9) | Full Transradial Prosthesis (Hand & Socket) | Functional performance (Brief Activity Measure for Upper Limb Amputees [BAM-ULA] score), user feedback | Yes |
| P4 | Development of a low-cost 3D-printed robotic prosthesis | Design & Development Report (n=1) | Full Transradial Prosthesis | Cost reduction, functional advantages/disadvantages | No |
| P5 | Affordable, open-source multifunctional prosthesis | Design & Development Report | Hand | Cost/time reduction, controller accuracy | No |
| P6 | Development of parametric prostheses for hand amputations | Design & Development Report | Hand | Personalized, functional, and accessible solution | No |
| P7 | Low-cost 3D-printed hand for children and distance-fitting methodology | Preliminary Study (n=11) | Hand | Feasibility of distance-fitting, user quality of life | Yes |
| P8 | Affordable anthropomorphic myoelectric prosthesis | Design & Development Report (n=10) | Hand | Cost reduction, user cognitive load (NASA Task Load Index [NASA-TLX]) | Yes |
| P9 | Parametric body-powered hand for children in developing countries | Design & Development Report | Hand | Parametric design feasibility, functional performance | No |
| P10 | AI-driven personalized myoelectric prosthesis | Primary Study (n=9) | Full Transradial Prosthesis (Hand & Socket) | Functional performance (BAM-ULA score), user feedback | No |
| P11 | Development of parametric prostheses for hand amputations | Design & Development Report | Hand | Personalized, functional, and accessible solution | No |
| P12 | Parametric 3D-printed body-powered hand | Design & Development Report | Hand | Improved customizability and usability features | No |
| P13 | Low-cost 3D-printed hand for children and distance-fitting methodology | Preliminary Study (n=11) | Hand | Feasibility of distance-fitting, user quality of life | Yes |
| P14 | Affordable, open-source multifunctional prosthesis | Design & Development Report | Hand | Cost/time reduction, controller accuracy | No |
| P15 | Parametric body-powered hand for children in developing countries | Design & Development Report | Hand | Parametric design feasibility, functional performance | No |
| P16 | 3D-printed non-metallic prosthetic arm for forequarter amputation | Case Study (n=1) | Full Forequarter Prosthesis | User satisfaction (Quebec User Evaluation of Satisfaction with Assistive Technology [QUEST 2.0]), cost reduction | No |
| P17 | Affordable anthropomorphic myoelectric prosthesis | Design & Development Report (n=10) | Hand | Cost reduction, user cognitive load (NASA-TLX) | Yes |
| P18 | Use of a 3D-printed transitional hand prosthesis for children | Observational Study (n=5) | Hand | Changes in anthropometry, Range of Motion (ROM), and strength | No |
| P19 | Bioinspired design of 3D-printed cellular metamaterial liners | Design & Simulation Study | Liner | Reduced contact pressure and deformation (Finite Element Analysis [FEA] results) | No |
| P20 | Design of a low-cost voluntary closing finger prosthetic | Case Study (n=1) | Finger | User feedback on fit, comfort, and functionality | No |
| P21 | Functional assessment of 3D-printed prostheses in children | Prospective Observational Study (n=14) | Hand | Functional dexterity (Box and Block test, Unilateral Below-Elbow Test [UBET]) | No |
| P22 | Digital reconstruction and mechanical prediction of prosthetics | Feasibility Study | General Prosthetic | Accuracy of digital reconstruction, prediction of mechanical properties | No |
| P23 | Functional 3D-printed finger prosthesis with compliant fingertips | Case Study (n=2) | Finger | Typing speed, comfort, dexterity, aesthetic satisfaction | No |
| P24 | Parametric body-powered hand for children in developing countries | Design & Development Report | Hand | Parametric design feasibility, functional performance | No |
| P25 | Design and evaluation of a functional and anthropomorphic hand prosthesis | Design & Evaluation Study (n=11) | Hand | Functional manipulation (Southampton Hand Assessment Procedure [SHAP]), user opinion on appearance | No |
| P26 | Affordable, open-source multifunctional prosthesis | Design & Development Report (n=3) | Hand | Grasping performance, controller accuracy | Yes |
| P27 | Parametric 3D modeling of a customized prosthetic finger | Design & Development Report (n=1) | Finger | Shape similarity, mechanism robustness (Finite Element Method [FEM]) | No |
| P28 | Accessible hand prostheses using 3D printing and smartphones | Thesis Summary | Hand & Socket | Feasibility of photo-based socket modeling, ease of implementation | No |
| P29 | Parametric production of prostheses using Multi Jet Fusion | Design & Development Report | Socket | Workflow efficiency, geometric accuracy | Yes |
| P30 | Interdisciplinary development of a customized 3D-printed prosthesis | Case Study (n=1) | Full Transradial Prosthesis | User comfort, ease of activation, satisfaction | Yes |
| P31 | Automated design and rapid manufacturing of customized prostheses | Design & Development Report | Full Upper Limb Prosthesis | Feasibility of automated design workflow | Yes |
| P32 | Overview of 3D-printed upper limb prosthetics | Literature Review | General Upper Limb | Synthesis of current solutions and challenges | No |
| P33 | Design of a 3D-printed hand with bio-inspired articulated fingers | Design & Development Report | Hand | Pinch force, mechanical performance | No |
| P34 | Review of 3D printing for upper extremity prostheses | Literature Review | General Upper Limb | Assessment of potential to address accessibility barriers | No |
| P35 | Review of technology, materials, and challenges of upper limb prostheses | Literature Review | General Upper Limb | Synthesis of factors affecting user suitability | No |
| P36 | Review of 3D-printed upper limb prostheses | Literature Review | General Upper Limb | Quantitative overview of existing devices and technologies | Yes |
| P37 | Review of open-source 3D printing in the prosthetic field | Literature Review | General Upper Limb | Overview of available open-source designs and efficacy data | Yes |
| P38 | Parametric design for customized transradial sockets | Design & Development Report | Socket | Design time optimization, preliminary user fit | No |
| P39 | Parametric design for customized transradial sockets | Design & Development Report | Socket | Design time optimization, preliminary user fit | No |
| P40 | Regression modeling for economical prosthetic limb prototyping | Modeling Study | General Limb | Development of a regression model for limb length estimation | Yes |

Methodology  Type:

- 23  studies  used  a  Design  &  Development  Report  or  similar  (including  design  &  simulation,  design  &  evaluation,  modeling,  thesis  summary).

- 12  studies  were  classified  as  Primary  Studies  (including  observational,  feasibility,  preliminary,  case  study,  and  prospective  designs).

- 5  studies  were  Literature  Reviews.

Prosthetic  Component:

- 17  studies  focused  on  the  Hand.

- 5  studies  focused  on  the  Socket.

- 5  studies  addressed  General  Upper  Limb  prosthetics.

- 4  studies  addressed  the  Full  Transradial  Prosthesis  (Hand  &  Socket).

- 3  studies  focused  on  the  Finger.

- One  study  each  addressed  the  following  components:

  - Hand  &  Socket

  - Full  Forequarter  Prosthesis

  - Full  Upper  Limb  Prosthesis

  - Liner

  - General  Limb

  - General  Prosthetic

Primary  Outcomes  (studies  may  be  counted  in  multiple  categories):

- Functional  performance:Dexterity,  manipulation,  ROM,  strength,  etc.,  was  a  primary  outcome  in  15  studies.

- Workflow  or  design  process:Efficiency,  feasibility,  accuracy  was  a  primary  outcome  in  10  studies.

- User  feedback:Satisfaction  or  quality  of  life  was  a  primary  outcome  in  9  studies.

- Customizability:Personalization  or  parametric  design  was  a  primary  outcome  in  8  studies.

- Cost  or  time  reduction:This  was  a  primary  outcome  in  6  studies.

- Controller  accuracy  or  cognitive  load:This  was  a  primary  outcome  in  5  studies.

- Mechanical  or  structural  performance:Force,  robustness,  FEA/FEM  was  a  primary  outcome  in  4  studies.

- Accessibility:Open-source  or  affordability  was  a  primary  outcome  in  4  studies.

- Preliminary  user  fit:This  was  reported  as  an  outcome  in  3  studies.

- Review  synthesis  or  overview:This  was  the  main  outcome  in  5  literature  reviews.

## Thematic Analysis

### Parametric Design Implementation and Anthropometric Integration

A  central  theme  is  the  use  of  parametric  design  in  Computer-Aided  Design  (CAD)  software  as  the  primary  mechanism  for  personalization  (P1,  P9,  P12,  P29).  In  this  approach,  a  base  digital  model's  key  dimensions  are  modified  by  user-defined  parameters  derived  from  anthropometric  data.  The  methods  for  collecting  this  data  reflect  a  trade-off  between  accuracy  and  accessibility.

Anthropometric  Data  Sources:

- High-fidelity  3D  scanning:The  most  frequently  cited  high-accuracy  data  source  is  3D  scanning  of  the  user's  residual  or  contralateral  limb  (P1,  P3,  P10,  P16,  P27,  P31).  This  provides  detailed  surface  geometry  for  creating  custom-fit  components,  especially  sockets.  For  instance,  two  studies  describe  workflows  that  begin  with  a  3D  scan  of  the  limb  or  a  mold,  which  is  then  used  to  generate  a  parametric  socket  model  (P1,  P29).

- Accessible  alternatives:Recognizing  the  barriers  of  specialized  scanning  equipment,  several  studies  explore  lower-cost  options.  One  study  pioneered  a  distance-fitting  method  using  measurements  from  2D  photographs  (P7,  P13).  More  recent  work  proposes  using  smartphone  video  for  digital  reconstruction  (P22)  or  combining  2D  photos  with  a  Statistical  Shape  Model  (SSM)  to  generate  a  3D  socket  design  (P28).  Manual  measurements  with  calipers  and  measuring  tapes  also  remain  a  viable  data  source  for  parameters  like  finger  lengths  and  palm  width  (P3,  P40).

The  integration  of  this  data  is  often  a  manual  or  semi-automated  process  of  inputting  measurements  into  a  CAD  program.  However,  a  trend  towards  more  automated  workflows  is  apparent.  Two  studies  developed  intelligent  or  algorithmic  CAD  models  that  adapt  automatically  based  on  3D  scan  data  (P29,  P31).

True  AI-driven  personalization  for  optimizing  fit  is  nascent.  One  study  presents  a  data-driven  framework  using  motion  and  pressure  data  to  optimize  socket  structure,  a  more  advanced  approach  than  simple  geometric  matching  (P2).  Other  studies  apply  machine  learning  not  for  physical  form  personalization,  but  for  classifying  Electromyography  (EMG)  signals  for  myoelectric  control  (P3,  P5),  predicting  mechanical  properties  of  printed  parts  (P22),  or  using  regression  modeling  to  estimate  limb  lengths  (P40).  The  evidence  shows  a  strong  reliance  on  parametric  modeling  driven  by  various  anthropometric  data,  with  AI  applications  emerging  more  in  adjacent  areas  like  control  and  material  science  than  in  the  direct  optimization  of  functional  fit.

### Functional Fit Improvements

Direct,  quantitative  comparisons  of  the  fit  achieved  through  parametric  personalization  versus  traditional  methods  are  exceptionally  rare  in  the  included  studies.  The  available  evidence  is  largely  qualitative,  based  on  preliminary  evaluations,  or  derived  from  simulations  rather  than  clinical  testing.

| Study ID | Fit Assessment Method | Improvement Metrics | Comparison Baseline |
| --- | --- | --- | --- |
| P1 | Preliminary user evaluation (n=1) | Qualitative fit and comfort assessment | Implied comparison to less efficient/customizable methods |
| P2 | User testing (n=5) with motion/pressure data | Improved adaptability, reduced muscle fatigue, higher satisfaction scores | Traditionally designed and fabricated socket |
| P3 | User feedback (n=9) | Qualitative reports of ease of use and comfort; notes on discomfort from sweating | High-end commercial devices (for performance, not fit) |
| P4 | Testing on a single patient | Easily scalable size | Traditional myoelectric prostheses (for cost) |
| P5 | No mention found | No mention found | No mention found |
| P6 | No mention found | No mention found | Implied comparison to conventional prostheses with acceptance issues |
| P7 | Comparison of photographic vs. direct measurements | No significant mean differences in measurements | Direct anthropometric measurements |
| P8 | No mention found | No mention found | Implied comparison to high-cost commercial prostheses |
| P9 | No mention found | No mention found | No direct comparison mentioned |
| P10 | No mention found | No mention found | Conventional myoelectric prostheses (for cost/adaptability) |
| P11 | No mention found | No mention found | Implied comparison to conventional prostheses with acceptance issues |
| P12 | No mention found | No mention found | Existing 3D-printed designs with limited customization |
| P13 | Comparison of photographic vs. direct measurements | No significant mean differences in measurements | Direct anthropometric measurements |
| P14 | No mention found | No mention found | No mention found |
| P15 | No mention found | No mention found | No direct comparison mentioned |
| P16 | User survey (QUEST 2.0) | Higher score for "accurate size" and "suspension" | User's current traditionally-made prosthesis |
| P17 | No mention found | No mention found | Implied comparison to high-cost commercial prostheses |
| P18 | No mention found | No mention found | No direct comparison made |
| P19 | Finite Element Analysis (FEA) simulation | Reduced peak contact pressure and deformation | Conventional silicone liners |
| P20 | User feedback (n=1) | Qualitative feedback on fit and comfort | "One shape fits all" open-source designs |
| P21 | No mention found | No mention found | Traditional cosmetic or body-powered prostheses (for function) |
| P22 | Digital reconstruction analysis | Average deviation < 2mm for digital model | Not applicable (methodology validation) |
| P23 | User evaluation (n=2) | Successful fitting reported | Performance with vs. without prosthesis |
| P24 | No mention found | No mention found | No direct comparison mentioned |
| P25 | No mention found | No mention found | No direct comparison made |
| P26 | No mention found | No mention found | Commercial prosthetics (for cost/complexity) |
| P27 | Deviation analysis | Similarity in shape between scan and 3D model | Not applicable (comparison to source scan) |
| P28 | Pilot study user testing | Error in automatic measurement was "too large for an acceptable socket design" | Manual measuring method |
| P29 | Surface correspondence analysis | Maximum surface difference of 0.93 mm (<1 mm tolerance) | Original composite laminated solution |
| P30 | User testing (n=1) | More appropriate and comfortable coupling with the limb | Initial prototype; implied vs. standard prostheses |
| P31 | Digital and physical testing | No mention found | UnLimbited Arm open source project |
| P32 | Not applicable (Review) | Not applicable | Not applicable |
| P33 | No mention found | No mention found | Conventional body-powered prostheses (for pinch force) |
| P34 | Not applicable (Review) | Not applicable | Not applicable |
| P35 | Not applicable (Review) | Not applicable | Not applicable |
| P36 | Not applicable (Review) | Not applicable | Not applicable |
| P37 | Not applicable (Review) | Not applicable | Not applicable |
| P38 | Preliminary user evaluation (n=1) | No mention found | No mention found |
| P39 | Preliminary user evaluation (n=1) | No mention found | No mention found |
| P40 | No mention found | No mention found | Implied comparison to standard sizing methods |

Fit  Assessment  Methods:

- We  found  user  testing,  evaluation,  or  feedback  as  the  fit  assessment  method  in  12  studies.

- We  found  digital  or  technical  analysis  (such  as  deviation  analysis,  FEA,  or  digital  reconstruction)  in  4  studies.

- We  found  measurement  comparison  methods  (e.g.,  photographic  vs.  direct  measurement)  in  2  studies.

- We  didn't  find  a  fit  assessment  method  in  17  of  the  35  non-review  studies.

Improvement  Metrics:

- We  found  qualitative  fit,  comfort,  or  usability  outcomes  in  8  studies.

- We  found  quantitative  fit  or  accuracy  outcomes  in  8  studies.

- We  didn't  find  improvement  metrics  in  21  of  the  35  non-review  studies.

Overall,  user-based  fit  assessment  and  qualitative  improvement  metrics  were  the  most  commonly  reported  approaches,  but  we  didn't  find  mention  of  fit  assessment  or  improvement  metrics  in  about  half  of  the  non-review  studies.

### User Comfort and Satisfaction Outcomes

User  comfort  and  satisfaction  are  frequently  stated  goals  but  are  often  assessed  through  informal  feedback  or  with  very  small  participant  groups.

Standardized  assessments:A  few  studies  employed  standardized  instruments.  One  study  with  nine  participants  used  the  Brief  Activity  Measure  for  Upper  Limb  Amputees  (BAM-ULA),  yielding  a  mean  score  of  7.4  out  of  10,  indicating  robust  functional  performance  (P3).  Participants  also  reported  ease  of  use,  though  some  noted  discomfort  from  sweating.  Another  study  used  the  QUEST  2.0  survey  with  a  single  user,  who  preferred  the  3D-printed  arm  for  its  ease  of  use  and  secure  suspension  compared  to  their  traditional  device  (P16).

Qualitative  feedback:Other  small-scale  studies  provide  further  insights.  One  study  reported  that  two  participants  experienced  improved  comfort,  reduced  pain,  and  aesthetic  satisfaction  with  custom  finger  prostheses  (P23).  Another  survey  of  11  children  and  their  parents  reported  that  a  low-cost  3D-printed  hand  had  a  significant  positive  impact  on  quality  of  life  and  daily  usage  (P7).

Reported  limitations:Negative  feedback  was  also  reported.  Low  durability  was  identified  as  a  significant  issue  in  two  studies  (P4,  P9),  and  another  noted  a  lower  durability  score  on  the  QUEST  2.0  survey  (P16).  A  functional  evaluation  with  14  children  found  that  while  3D-printed  prostheses  improved  gross  dexterity,  they  lacked  functionality  for  certain  pinch  grips  (P21).  This  suggests  that  while  personalization  can  improve  aspects  of  comfort  and  fit,  current  low-cost  designs  may  compromise  on  mechanical  performance  and  durability  compared  to  traditional,  more  robustly  manufactured  devices.  The  evidence  on  comfort  and  satisfaction  is  promising  but  preliminary,  limited  by  small  sample  sizes  and  a  lack  of  long-term  follow-up  in  most  studies.

### Manufacturing Efficiency and Customization Workflow

A  consistent  theme  is  that  digital  fabrication  workflows  using  parametric  design  and  3D  printing  improve  manufacturing  efficiency,  cost-effectiveness,  and  accessibility  compared  to  conventional  methods.  Numerous  studies  explicitly  aim  to  reduce  the  high  cost  and  long  waiting  times  of  traditional  prosthetics  (P1,  P4,  P5,  P8).

Cost  and  time  reduction:The  reported  cost  reductions  are  often  dramatic.

- One  study  estimated  the  cost  of  its  device  at  US$300  (P4).

- Another  reported  a  material  cost  of  only  €30  (P25).

- A  third  cited  a  material  cost  of  12  US  dollars  (P33).

- A  review  of  35  models  found  that  64%  cost  under  USD  50  (P34).  These  figures  contrast  sharply  with  the  thousands  of  dollars  for  conventional  devices  (P36).  Beyond  cost,  the  digital  workflow  streamlines  customization.  Several  studies  highlight  that  their  parametric  workflows  significantly  speed  up  the  design  phase  (P1,  P29),  with  one  focusing  on  a  fully  automated  process  from  3D  scan  to  a  printable  file  to  enable  mass  customization  (P31).  This  efficiency  is  particularly  impactful  in  low-resource  settings,  enabling  local,  on-demand  manufacturing  and  remote  fitting  (P7,  P28).

Workflow  challenges:Challenges  within  these  workflows  are  also  identified.

- One  study  noted  technical  limitations  in  CAD  software  that  constrained  their  parametric  model  (P1).

- Reviews  point  to  broader  systemic  challenges,  including  a  lack  of  standardized  manufacturing  protocols,  variability  in  design  quality,  and  issues  with  the  durability  and  mechanical  properties  of  common  3D-printing  materials  (P36,  P37).

These  challenges  can  hinder  the  translation  of  efficient  workflows  into  clinically  robust  and  reliable  devices.

# References

A. Bhat, Vamsikrishna M, and A. Acharya. “Functional Assessment of 3D Printed Prosthesis in Children with Congenital Hand Differences-A Prospective Observational Study.” The Journal of Hand Surgery (Asian-Pacific Volume), 2021.

A. I. Musa, Adebanjo Oluwakayode Sukanmi, and Folorunsho Uthman Oluwadamilare. “Utilizing Additive Manufacturing for Economical Prosthetic Limb Prototyping: A Guide from Regression Modelling.” Asian Health, Science and Technology Reports, 2024.

Alejandro M Saldarriaga, Enzo Romero, Victoria E Abarca, and Dante A Elias. “A Parametric Design Approach for Affordable Customized 3D Socket for Transradial Upper Limb Prostheses.” International Conference on Control, Decision and Information Technologies, 2024.

Alejandro M. Saldarriaga, Enzo Romero, Victoria E. Abarca, and Dante A. Elías. “A Parametric Design Approach for Affordable Customized 3D Socket for Transradial Upper Limb Prostheses.” International Conference on Control, Decision and Information Technologies, 2024.

———. “A Parametric Design Approach for Affordable Customized 3D Socket for Transradial Upper Limb Prostheses.” International Conference on Control, Decision and Information Technologies, 2024.

Enzo Romero, Jose G Garcia, Magno Par, Sebastian Caballa, Alejandro M Saldarria, Edson F Luque, Dante J Rodrigue, Victoria E Abarca, and Dante A Elias. “An Affordable AI-Driven and 3D-Printed Personalized Myoelectric Prosthesis: Design, Development, and Assessment.” IEEE Access, 2025.

Enzo Romero, Jose G. Garcia, Magno Parra, Sebastian Caballa, Alejandro M. Saldarriaga, Edson F. Luque, Dante J. Rodriguez, Victoria E. Abarca, and Dante A. Elías. “An Affordable AI-Driven and 3D-Printed Personalized Myoelectric Prosthesis: Design, Development, and Assessment.” IEEE Access, 2025.

Filip Gགྷrski, P. Zawadzki, Radosław Wichniarek, Wiesław Kuczko, Sandra Slupińska, and M. Zukowska. “Automated Design and Rapid Manufacturing of Low-Cost Customized Upper Limb Prostheses.” Journal of Physics: Conference Series, 2022.

I. Vujaklija, and D. Farina. “3D Printed Upper Limb Prosthetics.” Expert Review of Medical Devices, 2018.

J. Kate. “The FA3D Hand: Design and Evaluation of a Functional and Anthropomorphic Hand Prosthesis Using the Advantages of 3D-Printing,” 2016.

J. Lopez. “Accessible Hand Prostheses: 3D Printing Meet Smartphones,” 2021.

J. T. ten Kate, G. Smit, and P. Breedveld. “3D-Printed Upper Limb Prostheses: A Review.” Disability and Rehabilitation: Assistive Technology, 2017.

J. Zuniga, D. Katsavelis, Jean M Peck, John Stollberg, Marcin Petrykowski, Adam Carson, and Cristina Fernandez. “Cyborg Beast: A Low-Cost 3d-Printed Prosthetic Hand for Children with Upper-Limb Differences.” BMC Research Notes, 2015.

———. “Cyborg Beast: A Low-Cost 3d-Printed Prosthetic Hand for Children with Upper-Limb Differences.” BMC Research Notes, 2015.

J. Zuniga, Jean M Peck, Rakesh Srivastava, D. Katsavelis, and Adam Carson. “An Open Source 3D-Printed Transitional Hand Prosthesis for Children,” 2016.

Juan Sebastian Cuellar, D. Plettenburg, A. A. Zadpoor, P. Breedveld, and G. Smit. “Design of a 3D-Printed Hand Prosthesis Featuring Articulated Bio-Inspired Fingers.” Proceedings of the Institution of Mechanical Engineers. Part H, Journal of Engineering in Medicine, 2020.

Julio Fajardo, Victor Ferman, Ali Lemus, and E. Rohmer. “An Affordable Open-Source Multifunctional Upper-Limb Prosthesis with Intrinsic Actuation.” Workshop on Advanced Robotics and Its Social Impacts, 2017.

———. “An Affordable Open-Source Multifunctional Upper-Limb Prosthesis with Intrinsic Actuation.” Workshop on Advanced Robotics and Its Social Impacts, 2017.

Julio Fajardo, Victor Ferman, Ali Lemus, and Eric Rohmer. “An Affordable Open-Source Multifunctional Upper-Limb Prosthesis with Intrinsic Actuation.” Workshop on Advanced Robotics and Its Social Impacts, 2017.

Julio Fajardo, Victor Ferman, Diego Cardona, Guillermo Maldonado, Ali Lemus, and E. Rohmer. “Galileo Hand: An Anthropomorphic and Affordable Upper-Limb Prosthesis.” IEEE Access, 2020.

———. “Galileo Hand: An Anthropomorphic and Affordable Upper-Limb Prosthesis.” IEEE Access, 2020.

Junghun Lee, Chukwuemeka Nkama, Hadiza Yusuf, Joseph Maina, Jean Ikuzwe, Jean Byiringiro, Moise Busogi, and Conrad Tucker. “Accessible Digital Reconstruction and Mechanical Prediction of 3D-Printed Prosthetics.” Journal of Mechanical Design, 2025.

K. Ráž, Z. Chval, and Vladislav Kemka. “Parametric Production of Prostheses Using the Additive Polymer Manufacturing Technology Multi Jet Fusion.” Materials, 2024.

K. Wendo, O. Barbier, Xavier Bollen, T. Schubert, T. Lejeune, B. Raucent, and R. Olszewski. “Open-Source 3D Printing in the Prosthetic Field—The Case of Upper Limb Prostheses: A Review.” Machines, 2022.

Kelvin Loutan, and Umesh Persad. “The Design of a Low-Cost Voluntary Closing Finger Prosthetic for Developing Countries.” Conference on Computability in Europe, 2023.

Kendall F Gretsch, Henry D Lather, Kranti V. Peddada, C. Deeken, L. Wall, and C. Goldfarb. “Development of Novel 3D-Printed Robotic Prosthetic for Transradial Amputees.” Prosthetics and Orthotics International, 2016.

Lankha Wallace, Modar Hassan, Yukiyo Shimizu, Akira Kikuchi, and Kenji Suzuki. “Functional 3D-Printed Finger Prosthesis with Compliant Fingertips for Support of iADLs.” IEEE International Conference on Systems, Man and Cybernetics, 2024.

Letícia Alcará da Silva, F. O. Medola, O. V. Rodrigues, A. C. T. Rodrigues, and F. Sandnes. “Interdisciplinary-Based Development of User-Friendly Customized 3D Printed Upper Limb Prosthesis.” Advances in Usability, User Experience and Assistive Technology, 2018.

Ludovica Nini, Alessandro Ceccarelli, N. Tagliamonte, Loredana Zollo, and Fabrizio Taffoni. “Parametric 3D Modeling of a Customized Prosthetic Hand Finger for Additive Manufacturing.” International Conference on Biomedical Robotics and Biomechatronics, 2024.

M. Moreo. “Parametric Design of a 3 D Printable Hand Prosthesis for Children in Developing Countries,” 2016.

———. “Parametric Design of a 3D Printable Hand Prosthesis for Children in Developing Countries,” 2016.

———. “Parametric Design of a 3D Printable Hand Prosthesis for Children in Developing Countries,” 2016.

Marlene Bustamante, Rodrigo Vega-Centeno, Midori Sanchez, and Renato Mio. “[Regular Paper] A Parametric 3D-Printed Body-Powered Hand Prosthesis Based on the Four-Bar Linkage Mechanism.” International Conferences on Biological Information and Biomedical Engineering, 2018.

R. Brack, and E. Amalu. “A Review of Technology, Materials and R&D Challenges of Upper Limb Prosthesis for Improved User Suitability.” Journal of Orthopaedics, 2021.

Rodrigo Cézar da Silveira Romero, Kliftom Amorim Costa, Paulo Henrique Rodriguês Guilherme Reis, and Claysson Bruno Santos Vimieiro. “Development of Parametric Prostheses for Different Levels of Human Hand Amputations Manufactured Through Additive Manufacturing.” Applied Sciences, 2025.

———. “Development of Parametric Prostheses for Different Levels of Human Hand Amputations Manufactured Through Additive Manufacturing.” Applied Sciences, 2025.

Said Dababneh, N. Dababneh, Chenrui Xie, Hanna Henchi, and J. I. Efanov. “Shaping the Future of Upper Extremity Prostheses Through 3D Printing.” Prosthesis, 2025.

Trevor Binedell, Eugene Meng, and Karupppasamy Subburaj. “Design and Development of a Novel 3D-Printed Non-Metallic Self-Locking Prosthetic Arm for a Forequarter Amputation.” Prosthetics and Orthotics International, 2020.

Vasja Plesec, and Gregor Harih. “Bioinspired Design of 3D-Printed Cellular Metamaterial Prosthetic Liners for Enhanced Comfort and Stability.” Biomimetics, 2024.

Yu Gu, Long He, Haozhou Ze, Jiaxing Li, Ning Zhang, Xiufeng Zhang, Tao Liu, and Haozhou Zeng. “A Data-Driven Design Framework for Structural Optimization to Enhance Wearing Adaptability of Prosthetic Hands.” IEEE Transactions on Neural Systems and Rehabilitation Engineering, 2024.